## MODE
MODE
PACKET SIZE | DATA SIZE | DESCRIPTION | PC-CMV | VC-CMV | PC-SIMV | VC-SIMV | CPAP | BI-PAP | APR-VC | PSV | PC-CMV_BACKUP | VC-CMV_BACKUP
1 |  | HEADER:53 | HEADER:53 | HEADER:53 | HEADER:53 | HEADER:53 | HEADER:53 | HEADER:53 | HEADER:53 | HEADER:53 | HEADER:53 | HEADER:53
2 |  | HEADER:50 | HEADER:50 | HEADER:50 | HEADER:50 | HEADER:50 | HEADER:50 | HEADER:50 | HEADER:50 | HEADER:50 | HEADER:50 | HEADER:50
3 |  | DATA SIZE: 0x0D | DATA SIZE: 0x0D | DATA SIZE: 0x0D | DATA SIZE: 0x0D | DATA SIZE: 0x0D | DATA SIZE: 0x0D | DATA SIZE: 0x0D | DATA SIZE: 0x0D | DATA SIZE: 0x0D | DATA SIZE: 0x0D | DATA SIZE: 0x0D
4 | 1 | [STANDBY/RESUME]+MODE | 0x01 | 0x02 | 0x03 | 0x04 | 0x07 | 0x08 | 0x05 | 0x06 | 0x09 | 0x0A
5 | 2 | PIP | 0 -60 | NC | 0-60 | NC | 0-60[IPAP] | 0-60[IPAP] | 0-60[IPAP] | 0-60[IPAP] | 0 -60 | NC
6 | 3 | PEEP | 0-60 | 0-60 | 0-60 | 0-60 | 0-60[EPAP] | 0-60[EPAP] | 0-60[EPAP] | 0-60[EPAP] | 0-60 | 0-60
7 | 4 | VOLUME[LSB] | NC | 0 - 3000 | NC | 0 - 3000 | NC | NC | NC | NC | NC | 0 - 3000
8 | 5 | VOLUME[MSB]
9 | 6 | I[7-4 BITS]+ E[3-0 BITS] | [1-15] + [1+15] | [1-15] + [1+15] | [1-15] + [1+15] | [1-15] + [1+15] | [1-15] + [1+15] | [1-15] + [1+15] | [1-15] + [1+15] | [1-15] + [1+15] | [1-15] + [1+15] | [1-15] + [1+15]
10 | 7 | RR | [5-60] | [5-60] | [5-60] | [5-60] | [5-60] | [5-60] | [5-60] | [5-60] | [5-60] | [5-60]
11 | 8 | FIO2 | [0-7] | [0-7] | [0-7] | [0-7] | [0-7] | [0-7] | [0-7] | [0-7] | [0-7] | [0-7]
12 | 9 | RISE TIME[7-4]+TRIGGER TIME[3-0 BITS] | [0-10] | NC | [0-10] | NC | NC | NC | NC | NC | [0-10] | NC
13 | 10 | FLOW RATE | NC | [10-60] | NC | [10-60] | NC | NC | NC | NC | NC | [10-30]
14 | 11 | APNEA TIME [0-7 BITS] | NC | NC | NC | NC | [0-30] | [0-30] | [0-30] | [0-30] | NC | NC
15 | 12 | TRIGGER TYPE [7 BIT]+ TRIGGER LIMITS [3-0] | NC | NC | [0/1] + NC + [1-10] | [0/1] + NC + [1-10] | [0/1] + NC + [1-10] | [0/1] + NC + [1-10] | [0/1] + NC + [1-10] | [0/1] + NC + [1-10] | [0/1] + NC + [1-10] | [0/1] + NC + [1-10]
16 | 13 | T HIGH[7-4 BITS]+ T LOW[3-0 BITS] | NC | NC | NC | NC | NC | [1-15]+[1-15] | [1-15]+[1-15] | [1-15]+[1-15] | NC | NC
17 |  | CHECKSUM | 0xFF | 0xFF | 0xFF | 0xFF | 0xFF | 0xFF | 0xFF | 0xFF | 0xFF | 0xFF

## RANGE&ALERT
ALERT RANGE SET PARAMETER |  |  |  | ALERT
TOTAL PACKET SIZE | DATA SIZE | DESCRIPTION | DATA VLAUE& RANGE | TOTAL PACKET SIZE | DATA SIZE | DESCRIPTION | DATA VLAUE& RANGE
1 |  | H1 | HEADER:54 | 1 |  | H1 | HEADER:54
2 |  | H2 | HEADER:50 | 2 |  | H2 | HEADER:50
3 |  | LENGTH | DATA SIZE: 0x0D | 3 |  | LENGTH | DATA SIZE: 0x08
4 | 1 | _RANGE_MODE | 0x15 | 4 | 1 | HIGH_PRIORITY_MSB: _ALERT_MINT_VOLUME[ N/Y 0:1, L/H 0:1]+_ALERT_OXY[ N/Y 0:1, L/H 0:1]+_ALERT_PIP [ N/Y 0:1, L/H 0:1]+_ALERT_FLOW_ALERT[ N/Y 0:1, L/H 0:1] | 0b00000000
5 | 2 | _RANGE_VT_MIN | [0-250]x10 | 5 | 2 | HIGH_PRIORITY_LSB: _ALERT_APNEA[0:1]+ _ALERT_PATIENT_CIRCUIT[0:1] +_ALERT_OXYGEN SUPPLY[0:1]+_ALERT_BAT_DRAIN[0:1]+NC+NC+NC+NC | 0b00000000
6 | 3 | _RANGE_VT_MAX | [0-250]x10 | 6 | 3 | MEDIUM_PRIORITY_MSB: _ALERT_T_RR[ N/Y 0:1, L/H 0:1]+_ALERT_VT[ N/Y 0:1, L/H 0:1]+_ALERT_SPO2[ N/Y 0:1, L/H 0:1]+_ALERT_PULSE[ N/Y 0:1, L/H 0:1] | 0b00000000
7 | 4 | _RANGE_PIP_MIN | 0-60 | 7 | 4 | MEDIUM_PRIORITY_MSB: _ALERT_PEEP[ N/Y 0:1, L/H 0:1]+_ALERT_LEAK[0:1]+NC+NC+NC+NC+NC | 0b00000000
8 | 5 | _RANGE_PIP_MAX | 0-60 | 8 | 5 | LOW_PRIORITY: _ALERT_BLE[0:1]+_ALERT_MICRO_SD[0:1]+_ALERT_POWER[0:1]+_ALERT_NEBULIZER[0:1]+NC+NC+NC+NC | 0b00000000
9 | 6 | _RANGE_RR_MIN | 0-80 | 9 | 6 | NC BYTE | 0b00000000
10 | 7 | _RANGE_RR_MAX | 0-80 | 10 | 7 | NC BYTE | 0b00000000
11 | 8 | _RANGE_MINT_VOL_MIN | [0-120] | 11 | 8 | NC BYTE | 0b00000000
12 | 9 | _RANGE_MINT_VOL_MAX | [0-120] | 12 |  | CHECK SUM | 0xFF
13 | 10 | _RANGE_SPO2_MIN | [0-100]
14 | 11 | _RANGE_SPO2_MAX | [0-100]
15 | 12 | _RANGE_PULSE_MIN | [0-120]
16 | 13 | _RANGE_PULSE_MAX | [0-120]
17 |  | CHECKSUM | 0xFF

## Sheet1
SERVICE COMMAND |  |  |  | SERVICE REPORT PACKET
PACKET SIZE | DATA SIZE | DESCRIPTION | PC-CMV | TOTAL PACKET SIZE | DATA SIZE | DESCRIPTION | DATA VLAUE& RANGE
1 |  | H1 | HEADER:55 | 1 |  | H1 | HEADER:54
2 |  | H2 | HEADER:50 | 2 |  | H2 | HEADER:50
3 |  | LENGTH | DATA SIZE: 0x0D | 3 |  | LENGTH | DATA SIZE: 0x08
4 | 1 | MODE | 0x01 | 4 | 1 | _REPORT_DATA_CONNECTION[ 0/1]+_REPORT_BLOWER[ 0/1]+_REPORT_EXPVAL[ 0/1] +_REPORT_SENSOR [0/1]+_REPORT_ PROXIMAL[0/1]+_REPORT_SERVO[ 0/1]+_REPORT_LEAK[ 0/1]+_REPORT_ NEBULIZER[0/1] | 0b00000000
5 | 2 | _SERVICE_DATA_CONNECTION[0/1]+_SERVICE_BLOWER[0/1]+_SERVICE_EXP_VAL[0/1] +_SERVICE_SENSOR[0/1]+_SERVICE_PROXIMAL[0/1]+_SERVICE_SERVO[0/1] +_SERVICE_CIRCUIT_LEAK[0/1]+_SERVICE_NEBULIZER[0/1] | 0b00000000 | 5 | 2 | _REPORT_LED[ 0/1]+_REPORT_OXYGEN[ 0/1]+_REPORT_BATTERY[ 0/1] +_REPORT_I2C[ 0/1] | 0b00000000
6 | 3 | _SERVICE_LED[0/1]+_SERVICE_OXYGEN[0/1]+_SERVICE_BATTERY[0/1] +_SERVICE_I2C[0/1]+NC+NC+NC+NC | 0b0000XXXX | 6 | 3 | FLOW | 0b00000000
7 | 4 | NC | NC | 7 | 4 | PRESSURE | 0b00000000
8 | 5 | NC | NC | 8 | 5 | 02 FLOW | 0b00000000
9 | 6 | NC | NC | 9 | 6 | CIRUS | 0b00000000
10 | 7 | NC | NC | 10 | 7 | BATTERY | 0b00000000
11 | 8 | NC | NC | 11 | 8 | NC BYTE | 0b00000000
12 | 9 | NC | NC | 12 |  | CHECK SUM | 0xFF
13 | 10 | NC | NC
14 | 11 | NC | NC
15 | 12 | NC | NC
16 | 13 | NC | NC
17 |  | CHECKSUM | 0xFF

## Sheet3
CALIBRATION
BLOWER | DAC1 HIGH FLOW
LEAK | VOLUME HOLD
PRESSURE | TOP PRESSURE HOLD WITH LUNG AND WITHOUT LUNG
FLOW | DAC HIGH TOP FLOW TEST
O2 FLOW | DAC2 4.5v - O2 FLOW SENSOR 4.5V
BUZZER | HIGH LOW
BATTERY | HIGH LOW
PLAN: HEADER + MODE BYTE + NC [0xFF]+ CHECK SUM
RETURN : HEADER + 4&5 BYTE + NC [0xFF]+ CHECK SUM